## Instrument Errors
Instrument Error Tab
IE-1 --> IE-7, | IE-8 --> IE-12 | IE-14 | IE-16 | IE-17 --> IE-35, IE-37 --> IE-44, IE-46 | IE-36
Start test with 2 temp cycles | Start test with Non_Termination_Error asssay | Order valid test | Mobile Device connected to c360 | Postman; send error code | Postman; send error code
Postman; send error code | While test is running; verify error message appears | Insert cartridge without nfc tag | Postman; send error code | Start test, run to completion | Start test, run to completion
While test is running; Verify Error Message Appears | After test completion; tap error messages |  | Start test, run to completion | Verify Error Messages | Verify Error Messages
After Test Completion; Tap Error Messages | Verify Error Message Details |  | Verify Error Messages | Tap Error Message | Tap Error Message
Verify Error Message Details | Postman; get ../tests |  | Tap Error Message | Verify error details | Verify error details
Postman; get ../tests | Locate test and verify error code |  | Verify error details | Postman; ../module | Postman; ../tests
Locate test and verify error code |  |  | Search errors in c360 | Verify error code | Verify error code
Verify parameter match between postman and mobile device |  |  | Postman; ../module | Postman; send error code 0 | Postman; send error code 0
Postman; send error code 0 |  |  | Verify error code | Restart | Restart
Restart Omni Instrument |  |  | Postman; send error code 0 | Start a new test without errors | Start a new test without errors
Start a new test without errors |  |  | Restart |  | Omni Instrument not charging
|  |  | Start a new test without errors |  | Postman; ../module
Verify powerStatus & batteryStatus
View mobile device battery symbol
Connect charger to omni instrument
Postman; ../module
Verify powerStatus & batteryStatus
View mobile device battery symbol
Fatal Errors Tab
FE-1 --> FE-18 | FE-19 | FE-20-24 | FE-25 --> FE-27 | FE-28 --> FE-38 | FE-39 | FE-40 | FE-42 --> FE-51
Start test | Start test using valve defect failure assay | Start test with Non_Termination_Error asssay | Lost Communication | Postman; send error code | Corrupt DB | GX Term | Start test
Postman; send error code |  | After test completion; tap error messages |  | Start test, run to completion |  |  | Postman; send error code
After Test terminates; Tap Error Messages |  | Verify Error Message Details |  | Verify Error Messages |  |  | After Test terminates; Tap Error Messages
Verify Error Message Details |  | Verify parameters against Results worksheet |  | Tap Error Message |  |  | Verify Error Message Details
Postman; get ../tests |  |  |  | Verify error details |  |  | Postman; get ../tests
Locate test and verify error code |  |  |  | Postman; ../module |  |  | Locate test and verify error code
Verify parameter match between postman and mobile device |  |  |  | Verify error code |  |  | Postman; send error code 0
Postman; send error code 0 |  |  |  | Postman; send error code 0 |  |  | Restart Omni Instrument
Restart Omni Instrument |  |  |  | Restart |  |  | Start a new test without errors
Start a new test without errors |  |  |  | Start a new test without errors
Self-Test Errors Tab
ST-1 --> ST-21
Postman; send error code
Admin Actions; tap error
Verify error details
Restart Omni Instrument
Perform self-test
Start a new test without errors
Perform self-test
Miscellaneous Errors Tab
Whole Tab
Initialization Errors Tab
INITE-1 --> INITE-6, INITE-10 --> INITE-33 | INITE-7 | INITE-8 | INITE-9 | INITE-34 --> INITE-36 | INITE-37
cd /sandbox/home | Cartrde Tag Failed Validation | Available Space | No Assay for Product Code (2131) | Postman; send error code | No Power supply during FW Upgrade
vi InitError.txt |  |  |  | Start test, run to completion
error code |  |  |  | Verify Error Messages
dos2unix InitError.txt |  |  |  | Tap Error Message
Restart Omni Instrument |  |  |  | Verify error details
Verify Error Messsage & Details |  |  |  | Postman; ../tests
Verify error code
Postman; send error code 0
Restart
Start a new test without errors

## Test Results
Active Results Tab
Whole tab
Short test cases that require different assays.
View Results Tab
OMVR-1 | OMVR-2 | OMVR-3 | OMVR-4 | OMVR-5 | OMVR-6 | OMVR-7
Few test runs completed | SMS flag c360 | Test with error | Several tests with dates | Several test with info and dates | Several tests with dates and info | Retrieve omni tests in Dx
Tap on icons and verify them | Run test till completion | Click on test result detail screen | Run a few test till completion | Tap View Results | View Results and tap Search
| Tap on test in home screen | Verify test result detail screen | Tap View Results | Verify the table | Verify fields
| Take screenshot of patient info |  | Verify options | Verify horizontal drag | Enter info
| Tap on test in View Results |  | Tap filtering options & verify | Tap headers | Perform actions
| Verify test result details |  |  | Verify sort order
| Compare details against screenshot |  |  | Up and Down drags
| Long result and multi-line tests in view results |  |  | Select dats
| Verify multi line result |  |  | Verify tests
Verify long result
Test results having result color disclaimers
SCP instrument db
Open in SQLiteBrowser
Observe or modify result color disclaimers for tests
Replace db back in instrument
Verify colors in view results
Run a test and stop it
Verify NO RESULT
Run a test with INVALID result
Verify INVALID
Run tests with birthdate
Verify
Manage Results Tab
Whole tab
Requires disconnecting instruments and networks
Reporting and Printing Tab
Whole tab
Requires SMS and printing

## User Types
OM-US
OM-US-1
User Login and Logout
· Verify that username and password fields are displayed. Arrow to login is disabled until entry is made in username field.
· Verify that user shall only have capabilities to login or view information icons until a valid username and password are provided
· Verify field entry for login
Verify that a user shall be able to logout
OM-US-2
Change User’s own password (Test Admin, Guest and Admin User)
· Verify Admin accounts are unable to change their own password using Mobile Device.
· Verify Basic and Guest users are able to change their password.
· Verify Change password fields (Password, New Password, and Confirm Password). Arrow to change password is disabled until entries are made in each field.
· Boundary testing for password content
· Verify that user can login with new password and cannot login using old password
Verify password must be different than current password
OM-GU
OM-GU-1
Guest User Fields
· Verify Create Guest User fields (Email, Given Name, Family Name, ID Number, Password, Password Again) fields are displayed. Arrow to cancel is enabled and arrow to proceed is disabled.
· Boundary testing for Guest user fields
OM-GU-2
Canceling Guest User creation
· Verify guest user creation can be canceled
· Verify user data does not persist after canceling creation
OM-GU-3
Overwrite a Guest User definition
· Verify that creating an Admin or Mobile Device user with same username as guest user using C360 Admin will override the guest user definition.
· Verify test results are inherited by new Mobile Device user
OM-GU-4
Guest user creation is disabled
· Disabling of guest user is done through C360 Admin
· Verify Admin user is not able to create a guest user from the Mobile Device

## Ordering and Runnings Tests
OR-ORDR
OR-ORDR-1
Validation of Sample ID entry fields
Order entry conditions of the order entry fields: Sample ID, Test Type, Sample Type, Sample Type Dropdown list, Assay Name, Notes, Sample collection date and time
OR-ORDR-2
Validation of Patient ID entry fields
Additional fields will be tested through enabling in C360 Admin policies for displayed, required, optional.
Order entry conditions of the order entry fields: Patient ID, Address, Date of Birth, Given Name, Family Name, Gender, Race/Ethnicity.
OR-ORDR-3
Instrument assignment and test start
Verify the process/step header and icons enabled/disabled by completion of fields.
Verification of the workspaces for the process of ordering and running tests
· Start New Test
· Order entry fields
· Cartridge selection (NFC tag)
· Instrument selection/assignment
· Test start (Loading)
· Test running
· Test completion
OR-ORDR-4
Scan cartridge for a different assay
Select assay type in dropdown and then scan cartridge for a different assay. Verify message is displayed and user cannot start the test.
OR-ORDR-5
Attempt to select a different assay after scanning a cartridge
Scan a cartridge and then attempt to select a different assay in the dropdown. Verify that the user is not allowed to select a different assay that does not match the scanned cartridge.
OR-ORDR-6
Assay scanned does not match any installed assays
Scan a cartridge for an assay that is not installed on any connected instrument. Verify a message is displayed and user cannot start the test.
OR-ORDR-7
Turn off instrument while ordering a test
Turn off instrument during a test order. Tap the icon for the instrument and verify that a message is displayed and instrument cannot be selected.
OR-ORDR-8
Ordering a test when two mobile devices are connected
Issue a test order on one mobile device and attempt to begin the test on the associated instrument.
OR-ORDR-9
Attempt to order a test when battery is low
Connect mobile device to an instrument with low battery, battery below 20% charge, attempt to order test. Verify warning message is displayed and instrument is unable to begin test
OR-ORDR-10
Attempt to Order test when disk on instrument is full
Issue a test order on the associated instrument when said instruments disk is full, above 90%. Verify warning message is displayed and instrument is unable to begin test
OR-STAT
OR-STAT-1
Status display of Not Started, Running, and Completed test orders
Test Orders:
Verify the status for test orders/result that are pending/incomplete and processing/complete
OR-STAT-2
Omni instruments can’t be used to run a test
For the following use cases instruments cannot be used to run a test.
· Busy status
· Assay Upgrade
· Instrument SW Upgrade
· Self Test
OR-CAN
OR-CAN-1
Test Order Cancellation at Sample ID Option Screen
1) Cancel a test with all information entered and instrument selected but prior to placing the cartridge into the module
Module will ask for reset.
OR-CAN-2
Test Order Cancellation at Type Sample ID Information Screen
Test with no information entered
OR-CAN-3
Test Order Cancellation at Patient ID Information Screen
OR-CAN-4
Test Order Cancellation at Scan Cartridge Screen
OR-CAN-5
Test Order Cancellation at Scanned Cartridge Information Screen
OR-CAN-6
Test Order Cancellation at Selection of Omni Instrument Screen
OR-CAN-7
Test Order Cancellation after Omni Instrument is Selected Screen
OR-CAN-8
Test Order Cancellation after Test Start
Cancel test during test cartridge loading – ‘Loading’
Cancel test during test running
OR-TSEDIT
OR-TSEDIT-1
Edit fields after test start
· Verify that the user shall be allowed to edit test information at any time up until test completion.
· Verify that the logout process shall not affect any test being run by Omni instruments
OR-TSEDIT-2
Edit fields when test is completing
Verify that the user shall be prompted with a message stating test information cannot be edited since test has completed.
OR-TSEDIT-3
Edit test while running with multiple instruments
Verify test information edited on one instrument does not affect the information of test running on another instrument.
OR-BVAL
OR-BVAL-1
Format 3 NFC Tag
Verify Format 3 NFC tags are accepted
OR-BVAL-2
Format 4 NFC Tag
Verify Format 4 NFC tags are accepted
OR-BVAL-3
Format 5 NFC Tag
Verify Format 5 NFC tags are not accepted
OR-BVAL-4
Format 6 NFC Tag
Verify Format 6 NFC tags are accepted
OR-BVAL-5
Sample ID Barcode Formats
· Code 39
· Code 93
· Code 93 ASCII
· Code 128
· Code 128A
· Code 128B
· Code 128C
· Codabar
· Interleaved 2of 5
· Data Matrix
· QR Code

## Admin Actions
OM-INST
Om_Inst-1
Connect Omni Instrument to Omni Mobile Device
Om_Inst-2
Omni Instrument Details Screen on Omni Mobile device software
· Omni Instrument Nickname
o Default on connection – last 8 characters of S/N
o Editable, alphanumeric
o Must be between 1-8 characters
Om_Inst-3
Omni Instrument Actions
Om_Inst-4
Omni Instrument Maintenance
Om_Inst-5
Omni Instrument Self Test
Om_Inst-6
Omni Instrument status display
Om_Inst-7
Omni Instrument Battery
Low battery Pop up message
Om_Inst-8
Disconnect an Omni Instrument from a Lab
Om_Inst-9
Performing Omni actions on instrument associated with different lab
OM-DEVICE
Om_Device-1
Verification of the Mobile Device information page
Om_Device-2
Mobile Device Connection to Wi-Fi network
Om_Device-3
Users should get some visual feedback when action icons are tapped.
Navigating to each screen and verifying Visual feedback on all action icons.
OM-PRINTER
Whole tab

## Site Provisioning
HR
HR-1
· Generation of the one-time passwordGenerate a one-time password associated with the Institutional Administrator
· One-time password is sent via SMS to User’s on file mobile device
· One-time password is sent via valid email
HR-2
Attempt to login with user not known by C360 Admin at program start
HR-3
Attempt to login with Institutional admin user known to C360 Admin, valid password, valid OTP, no internet connection
HR-4
Attempt to login with Institutional admin user known to C360 Admin, bad password, valid OTP
HR-5
Attempt to login as Institutional Administrator known to C360 Admin, good password, invalid OTP
HR-6
Login as Cepheid Support user with valid password and valid OTP
HR-7
Login as Cepheid Admin user with valid password and valid OTP
HR-8
Login as Institutional Administrator known to C360 Admin, good password, valid OTP (perform setup).
· Institutional administrator logs into the Cepheid Omni Mobile Device software on the Mobile Device
· Select the site where the Omni Mobile Device Software is to be registered.
· The Mobile Device software then automatically obtains the institution key and performs initial site configuration synchronization with C360 Admin
· Verify Omni Mobile Device Software Device View screen after site provisioning is completed
Verify the Policy Data i.e. Guests allowed, Patient Demographics allowed/required, SMS destination
HR-9
Install Database viewing software and check the Omni Mobile Device database
· Check assays that are allowed by the current institution
· Handheld ID in both mobile device DB and C360 Admin are the same
Users that are associated to the site that the phone is provisioned to are the same.
HR-10
Manual Initialization process
Generation of unique application id.
OI
OI-1
Set up the Wi-Fi Network SSID on the Omni Mobile Device Software
OI-2
An Omni Instrument shall be associated with a Site.
The association is made in Omni Mobile Device Software and passed to C360 Admin.
OI-3
An Omni Instrument shall be re-assigned to a new Site.
The association is made in Omni Mobile Device Software and passed to C360 Admin.
OI-4
Install Database viewing software and check the Instrument database
· Check the instrument DB and C360 Admin list the same site
· Check the serial number is the same on both the C360 Admin and instrument DB
Users in both C360 Admin and instrument DB are the same
SCS
SCS-1
Site configuration upon software startup
· Set changes in C360 Admin to be downloaded to Omni Mobile Device Software. No network connection.
Set changes in C360 Admin to be downloaded to Omni Mobile Device Software. Network connection verified.
SCS-2
C360 admin auto sync after an hour since the last synchronization
SCS-3
User Changes Password in C360
Institutional Admin User passwords
SCS-4
Changing the Mobile Device to a new site
· Verify that information will change when sites are changed
Verify that when the site is changed the Omni Mobile Device Software will be reset to a state where it will only perform a site initialization
DOC
DOC-1
Verification of the documentation tab link connections for the Omni User’s Guide and Assays links
· Verification of the Omni Mobile Device Documentation Home screen
· Verification of the link to the Omni Mobile Device User’s Guide and internal links
Verification of the link to the Assays, each individual assay and internal links.
DOC-2
Site updates to the Assay and User Guide are downloaded and installed during site updates.
· Verify updates initiated in C360 Admin for the Omni User’s Guide are downloaded and shown in the documentation tab.
Verify updates initiated in C360 Admin for assay updates/additional assays are downloaded and shown in the documentation tab.
SP
SP-1 | ASSUMES OTP GENERATION IS POSSIBLE
Testing for the initial site provisioning.
Begin the setup process without a network connection.
SP-2
Testing for the initial site provisioning.
· Generation of the one-time password
· Associate a specific instance of the Omni Mobile Device software with an Institution Site.
Follow the setup process and ensure that the “Cancel Setup” option functions properly on the Network Configuration screen.
SP-3
Testing for the initial site provisioning.
· Generation of the one-time password
· Associate a specific instance of the Omni Mobile Device software with an Institution Site.
Follow the setup process and ensure that the “Cancel Setup” option functions properly on the Sync screen.
SP-4
Testing for the initial site provisioning.
· Generation of the one-time password
· Associate a specific instance of the Omni Mobile Device software with an Institution Site.
Follow the setup process and ensure that the “Complete Setup” option functions properly on the Sync screen.
SP-5
Testing for the initial site provisioning.
· Generation of the one-time password
· Associate a specific instance of the Omni Mobile Device software with an Institution Site.
Follow the setup process and ensure that the “Complete Setup” option functions properly on the Network Configuration screen without adding any Omni Instruments.
SP-6
Testing for the initial site provisioning.
· Generation of the one-time password
· Associate a specific instance of the Omni Mobile Device software with an Institution Site.
Follow the setup process to completion and ensure that the Network Configuration is working properly and a single Omni Instrument is connected and available.
SP-7
Testing for the initial site provisioning.
· Generation of the one-time password
· Associate a specific instance of the Omni Mobile Device software with an Institution Site.
Follow the setup process to completion and ensure that the Network Configuration is working properly and two or more Omni Instruments are connected and available.

## Handheld Events
HHE
HHE_1
User Management
· Add User
· Add Guest User
· User Login/Logout
· Guest/Handheld User changes password
· Admin changes Guest/Handheld User password
Admin user edits information about a Guest User
HHE_2
Instrument Management
Instrument Added
Instrument Removed
Instrument Restart
Instrument Shutdown
Instrument Available for Test
Instrument Unavailable for Test
HHE_3
Printer Connected (Bluetooth & WiFi)
HHE_4
Modify Network Settings
HHE_5
Test Run
Test is stopped
Test is cancelled
HHE_6
Instrument Maintenance
Weekly instrument maintenance started
Weekly instrument maintenance cancelled
Weekly instrument maintenance completed
Monthly instrument maintenance started
Monthly instrument maintenance cancelled
Monthly instrument maintenance completed
HHE_7
Synchronization with C360
HHE_8
Instrument Self-Test & Errors
Self-Test Started
Self-Test Completed
Self-Test Data
Instrument Error
Instrument Blocking Error
HHE_9
Remote Commands:
Added Assay to Instrument
Added Assay to Mobile Device
Upgraded Instrument Software
Command Accepted
Command Refused
HHE_10
Mobile Device Exception
HHE_11
Failed to Add Assay to Instrument
HHE_12
Command Error
HHE_13
Event IDs
Compare the event IDs to ensure that each ID is unique.

## Remote Commands
INTRODUCTORY ON AGILE

## Offline Mode
OFLRT
OFLRT-1
Offline Mode Mobile Device Screen
OFLRT-2
Running Tests in Offline Mode
· Ordering and running tests
o Add Sample to cartridge Step
o Completed Active and Inactive test result details
OFLRT-3
Run tests with multiple instruments connected
OFLRT-4
Print test result in Offline Mode
· Bluetooth Printer
· Wi-Fi Printer
OFLDTR
OFLDTR-1
Download test results to mobile device in offline mode
“Download test result from instrument” request
· Request not successful
· Request Successful
o Error conditions
o Recovery
OFLDTR-2
Test results download for Multiple instruments connected (offline mode)
· Request not successful
· Request Successful
· Error conditions
· Recovery
UPTR
UPTR-1
Offline to Online Mode Test Results upload with Omni instrument not connected
“Upload test results to C360 Server” request
· Request not successful:
· Request successful
o Error conditions
Recovery
UPTR-2
Offline to Online Mode Test Results upload with Omni instrument connected
“Upload test results to C360 Server” request
· Request not successful:
· Request successful
o Error conditions
o Recovery
o Verify after successful upload the test results are deleted from mobile device“rx” folder.
Verify “out” folder do not exist in mobile device
UPTR-3
Test results upload for Multiple instruments connected
UPTR-4
Uploaded test results information